## Slide 1: VIBEAITHON 2025
AI-Powered Smart Health
Monitoring and Early Disease Detection System
AURA-AI: Autonomous & Ubiquitous Remote-monitoring for
Advanced AI-powered Health

## Slide 2: The Problem: Bridging the Healthcare Gap
Despite advancements, accessible and continuous health monitoring remains a critical challenge, especially in remote and underserved areas. Early disease detection can save lives and reduce healthcare burdens.

## Slide 3: AURA-AI: Your Portable Health Companion
Our solution is a compact, AI-powered health box designed for Autonomous & Ubiquitous health monitoring.
ECG
AD8232 sensor for heart
rhythms
SpO2
MAX30102 oxygen and pulse
Air Quality
MQ135 detects pollutants
Temp & BP
Skin temp and cuff-less PTT

## Slide 4: INNOVATION & EDGE: WHY AURA-AI STANDS OUT
Cleaner ECG Signals
Advanced noise reduction for precise cardiac data.
On-Device Cardiac Detection
Real-time analysis at the edge for immediate insights.
Cuff-less Blood Pressure
Comfortable, continuous BP monitoring without a cuff.
Smarter Breathing Check
Advanced analysis for respiratory patterns and health.
Vitals Forecasting
Predictive health analytics and early risk assessment.
Personalized Risk Index
A comprehensive, dynamic assessment of individual health risks.

## Slide 5: TECHNICAL APPROACH: FROM SENSOR TO CLOUD
Methodology Flow
Sensor Fusion
Collect ECG, SpO₂, PPG, respiration (accelerometer), temperature, and environmental data simultaneously.
Signal Processing
Apply real-time filters, QRS detection, and PTT calculation for clean ECG and cuff-less BP trends.
Multi-Parameter
Analysis
Combine HR, SpO₂, BP, respiration, and environment → generate health status insights.
User Interface
Interactive dashboard with ECG waveform, BP trend, SpO₂, and respiration graphs in real time.
Cloud & Alerts
Auto-upload to ThingSpeak; instant alerts sent to caregivers/doctors via Telegram during abnormalities.
Key Components
Hardware
ESP32 & Raspberry Pi
AD8232 (ECG)
MAX30102 (SpO₂)
MQ135 (Air Quality)
Software
Python, TensorFlow Lite
ThingSpeak (Cloud)
Telegram API (Alerts)

## Slide 6: CHALLENGES & MITIGATION
Our team successfully navigated several technical hurdles to ensure AURA-AI's reliability and precision.
MAX30102 SpO₂ Module Errors
Initial MAX30102 SpO₂ modules purchased in India had PCB errors. This led to inaccurate
readings and instability.
Fix: We performed board-level soldering to correct the traces and implemented careful
calibration procedures to ensure accurate SpO₂ readings.
AD8232 ECG Electromagnetic Interference
The AD8232 ECG modules experienced significant electromagnetic interference, causing noisy and unreliable signals.
Fix: We employed advanced digital filtering techniques and ensured proper grounding of
the circuitry to eliminate interference, leading to cleaner ECG signals.
SpO₂ Noise Spikes
Sporadic noise spikes in SpO₂ readings disrupted data consistency, making continuous
monitoring challenging.
Fix: Median filtering was applied to the SpO₂ data stream, effectively smoothing out
transient noise spikes without losing critical information.
Hardware Validation
Ensuring the accuracy and reliability of sensor data required robust hardware validation against known standards.
Fix: Validation was performed using accelerometer swing tests for motion artifacts and
SpO₂ skin resistance tests for precise calibration, ensuring data integrity.
Optimized Sampling Rate
An unoptimized sampling rate could lead to either data overload or loss, impacting cloud
integration and real-time analysis.
Fix: We meticulously optimized the sampling rate to balance data fidelity with efficient
and stable cloud integration, preventing data loss.
BP Accuracy (PTT-based)
Achieving reliable cuff-less blood pressure measurements based on Pulse Transit Time (PTT) presented an accuracy challenge.
Fix: The PTT-based system now provides continuous trend monitoring, offering insights into BP changes, with optional cuff calibration for personalized baseline accuracy.

## Slide 7
Desirability
Target users: Rural clinics, ambulances, elderly-care homes, and home-based patients.
Pain point solved: Eliminates need for multiple costly instruments and continuous manual monitoring
User validation: Early prototypes tested with biomedical datasets and live signals show reliable ECG and SpO₂ trends; doctors confirm usability for primary-care screening.

## Slide 8: FEASIBILITY & VIABILITY: A PRACTICAL SOLUTION
FEASIBILITY:
It uses low-cost, widely available modules (ESP32, AD8232, MAX30102, accelerometer, Pi 4)
Built on open-source frameworks (Arduino, Python, TensorFlow Lite, ThingSpeak, Telegram API)
VIABILITY:
Prototype-ready: Hardware assembled, tested with datasets + real signals.
Future-ready: Strong path to TRL upgrade, medical-grade sensors, and
regulatory approvals (ISO/CE).
DEPLOYMENT ENVIRONMENTS:
Our system is designed for broad applicability:
Seamless integration into clinics, emergency services, and home care.
ECONOMIC VALUE:
Affordable alternative to expensive bedside monitors.
Portable design makes it deployable in homes, ambulances, and mobile health units
Scalable platform: Expandable to Cloud Dashboards,
telemedicine networks & wearables.
Supports National Digital Health Mission and aligns with UN SDG Goal 3
(Good Health & Well-being).

## Slide 9: IMPACT & BENEFITS: TRANSFORMING HEALTHCARE
Rural Patients
Elderly
Chronic Illness
Emergency Services
Targeting diverse segments for maximum health impact.
Clinical Impacts
Early Risk Detection
Proactive intervention for critical conditions.
Continuous Monitoring
24/7 health oversight for peace of mind.
Data-Driven Insights
Personalized healthcare through intelligent analytics.

## Slide 10: Viability
Revenue / Impact Model: Low-cost subscription or one-time device sale for clinics and home users.
Market Potential: Rising demand for portable health diagnostics and telemedicine integration across India.
Scalability: Modular design allows future wearable version; supports integration with NDHM / Aarogya Setu Health Stack.
Cost: Estimated below ₹2,500 per unit in mass production.

## Slide 11: Prototype

## Slide 12: Prototype

## Slide 13: Prototype

## Slide 14: Prototype
Working features video
LINK: https://drive.google.com/file/d/16Ad1lGwIH42pDH1Y6s1ycDH8BzEPEF2q/view?usp=sharing

## Slide 15: Prototype
Demonstration video
LINK: https://drive.google.com/file/d/1oahr23VaH1ktzGe77dklZe2Q2VhNCi7k/view?usp=drive_link

## Slide 17
Team & Roles
Name: Abisek Sasikumar
Role: Solo Developer & Researcher
Skills & Expertise:
Embedded Systems (ESP32, Raspberry Pi)
Biomedical Sensor Integration (ECG, PPG, MQ sensors)
Signal Processing (ECG filtering, PTT estimation)
Machine Learning & Deep Learning (TensorFlow Lite, 1D-CNN, LSTM)
Cloud & IoT Integration (ThingSpeak, Telegram API)
Data Visualization (Python, PyQtGraph, Matplotlib)
Why I Can Complete This Project: With a background in Electronics, Sensors, and AI-based Health Monitoring, I have hands-on experience building and testing multi-sensor biomedical systems and implementing real-time AI inference on embedded platforms. The complete prototype is already functional and ready for further optimization.

## Slide 18
ASK
Mentorship: AI model validation and medical data calibration support.
Partnerships: Collaboration with biomedical research labs or hospitals
.
Funding: For hardware optimization and wearable-scale production.

## Slide 19: RESEARCH & DEVELOPMENT
Key References & Resources:
ECG & PQRST Analysis: MIT-BIH Arrhythmia Database (ECG datasets)
Cuff-less Blood Pressure (PTT): Blood Pressure Estimation Based on Photoplethysmography for Personalized Healthcare
Respiration Monitoring (PPG + Accelerometer):
  "Respiratory Rate Monitoring Using PPG Signals" – Journal of Clinical Monitoring & Computing
  ADXL345 datasheet & biomedical motion studies
Environmental & Multi-Sensor Fusion:
  Bosch BMP280 Environmental Sensor White Paper
  •MQ-135 Gas Sensor for Air Quality Studies
Explore our progress and code:
HEARTWARE

## Slide 20: Abisek S
Linkedin:
https://www.linkedin.com/in/abisek-sasikumar-42a89327a/
Email:abiseksasikumar@gmail.com
Mobile No:9080991640

## Slide 21: Thank You
Together, we can build a healthier India.